| DESIGN REVIEW REPORT | Report No. TR-391-003-30-0 |
| --- | --- |

| The Design Review Report Shall include at a minimum: |
| --- |
| The title of the item or system; A description of the item; Design Review Report Number; The type of design review; The date of the review; The names of the presenters The names, institutions and department of the reviewers The names of all the attendees | Findings/List of Action Items – these are items that require formal action and closure in writing for the review to be approved. See SLAC Document AP-391-000-59 for LUSI Design Review Guidelines. Concerns – these are comments that require action by the design/engineering team, but a response is not required to approve the review Observations – these are general comments and require no response |
| TYPE OF REVIEW: Advanced Procurement Review |
| WBS: |

| Title of the Review | XCS Diffractometer, Large Angle Detector Mover and Large Offset Monochromator Advanced Procurement Technical Review |
| --- | --- |
| Presented By: | Aymeric Robert |
| Report Prepared By: | Soon-Hong Lee(ANL) | Date: 23 April 2009 |
| Reviewers/Lab : | Soon-Hong Lee(ANL), Tom Rabedeau (SLAC), Mohan Ramanathan (ANL), Deming Shu (ANL) |
| Distribution: |  |

| Attachments: | Review Slides Design Checklist Calculations Other |
| --- | --- |

# Purpose/Goal of the Review:

# Provide a technical review of the procurement specifications, statements of work and procurement strategy for the XCS Diffractometer, Large Angle Detector Mover System and Large Offset Monochromator.

# Attendees:

Eric. Bong, David Fritz, Tom Fornek, Hanibal Joma

XCS Diffractometer/Large Angle Detector Mover/Large Offset Monochromator          Advance Procurement Technical Review

# Date: April 23, 2009

# Review Committee

Soon-Hong Lee(ANL), Tom Rabedeau (SLAC),

Mohan Ramanathan (ANL), Deming Shu (ANL)

# Review Charge

DOE-BES mandated review to assess the advance procurement readiness of the X-ray Correlation Spectroscopy Science (XCS) Diffractometer, Large Angle Detector Mover, and Large Offset Monochromator.

- Review the specifications and determine whether the specifications meet the overall requirements and are reasonable and achievable given current technology.

- Evaluate the list of potential vendors with whom the team has had discussions.

- Review the acquisition strategy for the systems and determine if the strategy is sound and will allow timely delivery of the full system.

- Review the vendor selection criteria and determine if these criteria and the quality assurance plan are appropriate for the system.

# Presenter

Aymeric Robert

# Summary

The purpose of this review is to review the specifications, the procurement process, and the potential vendors of the long-lead items of the XCS instruments since the LUSI project has not reached CD-3 and DOE wishes to review all the procurement specifications to make sure that LUSI team effectively uses DOE funding. For the committee to examine the associated documents as shown in Table, Aymeric Robert presented the XCS project overview, followed by a general discussion of performance requirements of the diffractometer, large angle detector mover, and large offset monochromator.

| Doc. No. | Title | Remarks |
| --- | --- | --- |
| PS-391-001-70 | LUSI X-Ray Correlation Spectroscopy Instrument Diffractometer System | Diffractometer |
| PS-391-001-71 | Bid Evaluation Guide for the Procurement of the XCS Diffractometer System |  |
| SP-391-001-30 R0 | LUSI XCS Diffractometer System Engineering Specification Document |  |
| SP-391-001-32 R0 | Physics Requirements for the XCS Diffractometer System |  |
| PS-391-001-75 | LUSI XCS Large Angle Detector Mover Main Carrier Mover Subsystem - Procurement Specification | Large Angle Detector Mover |
| SP-391-001-31 R0 | LUSI XCS Large Angle Detector Mover Engineering Specification |  |
| SP-391-001-33 R1 | Physics Requirements Document for the XCS Large Angle Detector Stage |  |
| SP-391-001-77 R0 | XCS Large Angle Detector Mover Main Carrier Mover - Statement of Work |  |
| PS-391-001-78 | LUSI XCS Large Offset Monochromator -Procurement Specification | Large Offset Monochromator |
| PS-391-001-79 | LUSI DCO Large Offset Monochromator Bid Evaluation Guide |  |
| SP-391-00-16 R0 | Physics Requirements for the LUSI Large Offset Monochromator |  |
| SP-391-000-94 R0 | LUSI Large Offset Monochromator Engineering Specification |  |
| SP-391-001-80 R0 | XCS Large Angle offset Monochromator - Statement of Work |  |

In general, the specifications and procurement documents including the acquisition strategy and the vendor selection criteria for XCS diffractometer, large angle detector mover, and large offset monochromator are technically sound. But, there are some technical design & interface issues that should be addressed before proceeding for advanced procurement.

The observations, findings, and recommendations of each instrument are described in more detail in the following sections.

# A. XCS Diffractometer

A.1 Observations:

- Advance procurement prior to CD3 is required in order to meet early science schedule.

- The demanding XCS diffractometer is a 4-circle horizontal scattering geometry diffractometer to position the sample to the x-ray beam in a precise manner.

- The total weight of the local detector will be of the order of 10 kg, maximum. The local detector is not part of the scope of the diffractometer.

- The operating envelope of the sample is demanding with +3.0 ~ -30.0 mm in the vertical direction and +2.0 ~ -2.0 mm in the horizontal direction at each Interaction Point (IP). The distance between the 1st IP and the 2nd IP is of 600.0 mm.

- The nominal operation height of the diffractometer is 1,400.0 mm from the floor. The sample manipulation stage should accommodate a 50.0 kg (110.0 lb) payload with a center of mass up to ~50.0 mm (2.0 inch) above the goniometer center of rotation.

- Movable diffractometer base will be used to completely move the diffractometer from/to 3 different positions (IP1, IP2, Park position). The demanding base consists of solid granite with guiding rail on the side, multiple air-pads, and floor insert. The required position repeatability is specified as less than ± 0.5 mm. This motion is not motorized.

- The sphere of confusion of the diffractometer is defined better than 70 μm. The experimental sphere of confusion including various degrees of freedom of the local detector is defined as not exceed 100 μm.

- The position drift of the sample is defined by less than 10 μm in X, Y, and Z, and less than 0.3 mDeg (~1 arcsec) in roll, pitch, and yaw over 1 day.

- The engineering specification document (ESD) specifies that the center of rotation vibration is less than 1 μm and vibration does not cause CoR angular deviations greater than 0.06mDeg (~0.2 arcsec).

A.2 Findings:

- The sample goniometer system is in a sufficiently developed state though there are some technical design issues noted below that should be addressed before procurement.

  - Specification of the performance of each stage in isolation does not properly account the errors introduced by coupling multiple stages. A more robust specification would tolerance the degrees of freedom at the sample interaction point.

  - By the same token, the specification should define the sample positioning needs and performance envelope not the engineering solution. The specification should allow the vendor as much freedom of engineering as feasible consistent with meeting these requirements. Thus, for example, the YY translation stage need not be specified as interleaved between the theta and chi/phi stages.

  - Cross coupling problem of the vertical stage YY to the tilting stages of χ and φ is anticipated. As noted above in the same context, stack-up position of these stages should be re-visited. The vendor may have better capability to build and select the stage stack-up position. The order of the stacking of the various motions should not be dictated, but left to the vendor who is designing to specifications. So, the ESD should specify that the vendor should propose their capability on the order of stack-up of the stages to achieve the motion stability requirements.

  - The technical specifications for stage YY should address the motion straightness of trajectory.

- The XX_air motion is a good idea to move the detector. However, it would be advisable to make the table rest in fixed points on the floor. It would be advisable to think about installation cups recessed in the floor to provide repeatability (+/- 0.5 mm) of the diffractometer position after motion. In addition, the number of air-pads between the base and the floor will affect the overall system dynamics (high natural frequencies) and positioning stability. To meet the position drift and vibration requirements specified in the ESD, the interface design should be re-visited to ensure overall positioning stability and make the device more robust to the floor vibration.

- The ESD does not include its minimum warranty period. Thus, for example, the ESD may include “the warranty will start the date of approval of its final acceptance test from SLAC for a period of one year“. The procurement document includes only “Proposed guarantee“.

- Synthetic granite material may be considered as an option due to its excellent damping quality.

- “True Stone Distributing” was a supplier of the granite in the APS nanoprobe project. The LUSI may consider this company as a subcontractor to the vendor.

- There is a typo in Section 7.8.2 of the ESD.

- The vendor would benefit from as much hutch environmental information as available.

- General procurement plan is appropriately documented though the "not acceptable" classification in the bid evaluation matrix should be removed. If the vendor offers an unacceptable solution, then the vendor's proposal should be disqualified not down rated.

A.3 Recommendations:

- The specification should allow the vendor as much freedom of engineering as feasible consistent with meeting the performance requirements.

- The XCS diffractometer proceeds to advance procurement with minor changes to the specifications.

# B. XCS Large Angle Detector Mover

B.1 Observations:

- Advance procurement prior to CD3 is required in order to meet early science schedule.

- The demanding XCS Large Angle Detector Mover (LADM) consists of the main mover carriage, an adjustable end module, and a modular vacuum system to prevent air scattering. The end module further consists of a beam-stop module with an exit window, a detector precision positioning stage, and a local diode detector to aid its alignment.

- The overall motion of the LADM to allow placing the XCS detector about a truncated sphere is of approximately 8.0 m maximum radius (with also the possibility of a shorter approximately 4.0 m sample-detector distance) with the center coincident with the experimental interaction point.

- The LADM should have the capability of changing its Center or Rotation (horizontally and vertically), in order to properly adjust for any change in the experimental interaction point. The main mover carriage should rotate in the horizontal plane and cover scattering angles ranging from +55.0° to -1.0° with desired resolution of 1.0 mDeg. The main mover carriage should also provide the capability to tilt in the vertical plane over a limited range of angles of -0.1° to 1.0° with desired resolution of 1.0 mDeg. Due to the pivoting motion of this tiling and the vertical motion of -30.0 mm to +3.0 mm of the diffractometer, the overall range of vertical translation at the downstream unit is at a minimum -45.0 mm to +145.0 mm. Its maximum load capacity is ~1,814.3 kg (4,000 lbs).

- Controls will be provided from SLAC (will be obtained from APS-HERIX).

- In the vacuum chamber system, CVD diamond is specified as the entrance window material. Its diameter is 40.0 mm and it could be replaced with Kapton.

- The exit window system is located right after the in-vacuum beam-stop. A thin foil of Kapton is specified as its material.

B.2 Findings:

- It is difficult to imagine that this system warrants award to an offshore vendor. Surely there are US vendors for machine hardware that could meet the specification. The plan to separate the detector motion assembly and the flight tubes from the main mover is a good idea. This will help in expediting the main detector mover design and procurement, while the details of the flight path and the detector table designs are completed in house. The division of effort envisioned whereby SLAC staff takes responsibility for the flight tube, beam stop, etc is appropriate.

- The stored energy associated with the evacuated drift tube is substantial. The ES&H issues associated with this stored energy in the presence of a thin window should be addressed early in the design effort. The use of a large area, thin Beryllium window should be re-evaluated in this context. A thin foil of Kapton specified in the ESD is preferred.

- The large angle detector mover design is based on one currently in use at APS in sector 30. While there are many similarities, some of the differences are significant to affect the design for the LUSI instrument.

- The load of the detector and the associated flight path is significantly small compared to the APS device.  This has the advantage of reduction in cost and will provide better stability for the mover.

- Renishaw’s linear encoder used for HERIX at Sector 30, APS, has a reliability problem so other options to measure the displacement should be considered. Laser interferometer is an option to be evaluated.

- Hardware controlled speed limit is critical for the safe operation of a detector mover with large travel range.

- X-ray sensors may need to be added to the beam stop to provide an interlock for the detector equipment safety.

- In the physics document (SP-391-001-33), the Large Angle Detector Stage (LADS) should be the Large Angle Detector Mover (LADM).

- There is a typo in Section 4 in the ESD (SP-391-001-31). The reference document should be SP-391-001-33.

- In Section 4.1 of the procurement specification (PS-391-001-75), “Diffractometer system” should be “Main Carrier Mover subsystem of the Larger Angle Detector Mover”. In Section 10-d) of the procurement specification, “SP-391-001-30” should be “SP-391-001-31”.

- General procurement plan is appropriate though the "not acceptable" classification in the bid evaluation matrix should be removed. If the vendor offers an unacceptable solution, then the vendor's proposal should be disqualified not down rated.

- The vendor would benefit from as much hutch environmental information as available.

B.3 Recommendations

- It is recommended that the large angle detector mover carriage should proceed to advance procurement after defining the interface between the in-house design and the mover carriage.

- The large Beryllium window for the detection drift tube has serious consequences for the ES&H if it ruptures. So, a thin foil of Kapton is recommended for this application.

# C. XCS Large Offset Monochromator

C.1 Observations:

- The demanding Large Offset Monochromator (LOM) is to deflect the x-ray beam 600 mm horizontally parallel to the incoming beam and to select specific photon energy from 4 keV to 25 keV.

- Instead of placing the first crystal in a fixed position and moving the second crystal to cover the required energy range like a Sector 12’s monochromator currently built at the APS, both crystals will be moved half the range of the required longitudinal translation like a monochromator built at DESY for PETRA III.

- Two pairs of crystals, Si (111) and Si (220), will be used in the LOM and diamond (111) and (220) are considered as a possible upgrade.

- The FEH hutches and utilities are not built yet. After completion of the facilities in the end of this year, vibration spectrum can be supplied to the vendor.

C.2 Findings

- The specification for the large offset monochromator for LUSI is based on APS design. While there are major similarities between the two designs, care has to be taken to revise the specifications for the LUSI instrument.

- The plan to use a translation stage in the Z axis along the beam direction for both first and second crystals is an excellent match for the LCLS source.  The design as called for is a mirror image between the first and second crystal motions.  The design also has the advantage of providing a fixed reference point for a potential detector between the two crystals. This monochromator concept of moving both crystals symmetrically about a fixed point is quite advantageous given the absence of cooling required of the first crystal. This greatly eases the Z rail tolerance problem.

- The LUSI staff is planning to design the crystal mounts and will be trying out different crystals during the course of their commissioning. Based on the past experience at SSRL and at APS, it is advisable to define a clear interface point for the crystal mount.

- A number of the fine motion degrees of freedom are rightfully the domain of the crystal mount engineering and should be extracted from the vendor statement of work. The vendor SOW would benefit from reduced scope to include only the two larger translations and large angle rotation for each crystal. The SOW should be clear that the vendor needs to provide the means to swap crystals. Since vendors may not be thinking outside the box, the SOW might include descriptions of several means to swap crystals (e.g., vertical translation, 180 deg rotation and horizontal translation, ...)

- As noted above in a different context, the specification should allow the vendor as much freedom of engineering as feasible consistent with meeting the monochromator requirements.

- Entropy is real, so the monochromator stability is likely to be enhanced by including the minimum number of degrees of freedom necessary. Also the lighter the load carried by a degree of freedom, the higher the resonant frequency. Thus fine motions are best implemented in the crystal mount.

- As noted above in the same context, the removal of some of the motions to the crystal mounts will help in expediting the main monochromator procurement. Every extra degree of freedom is a source of operational instability and will degrade the performance of the device. Based on experience it would be advisable to remove the H1, H2,  and  advisable motions from the main stage and make them part of the crystal mount.  These designs exist at SSRL and can be reused in minimum work.  The experience of the SSRL engineering staff should be valuable in the design of the crystal mount.  The interface between the crystal mount and the rest of the stages should be clearly defined in the ESD and its crystal mount design effort by SLAC engineer should be included in budget.

- Some crystal motions, such as Y1, Y2, etc., could be covered by a SLAC-home-made crystal holder as well.

- Care has to taken to clearly define the tolerances on all the stages and ask for a testing plan of the deliverables.

- Radiation hardened optical encoders may be necessary for encoders inside the vacuum chamber.

- Support structures of the monochromator should have overall alignment capability (wedge Jack and X and Y pusher etc.)

- The vendor would benefit from as much hutch environmental information as available.

- “Angle” should be removed from a Title of “XCS Large Angle offset Monochromator Statement of Work – Statement of Work” (SP-391-001-80 R0)

- In Section 3.2.2. of procurement document (PS-391-001-78), “SP-3914-000-94“ should be “SP-391-000-94”. In Section 3.2.3. “Main Carrier Mover subsystem” should be “Large Offset Monochromator”.

- A possible US vendor for the monochromator is as follows;

Physical Sciences Laboratory

University of Wisconsin-Madison

3725 Schneider Dr. Stoughton WI 53589-3098

Phone:	(608)877-2200, Fax: (608)877-2201

Email: ffeyzi@psl.wisc.edu

Contact: Farshid Feyzi

- The project should plan for a 24 month lead time following order placement. The vendor may deliver in 18 months and the crystal mounts should be engineered accordingly, but the definition of success should be 24 months.

- General procurement plan is appropriate though the "not acceptable" classification in the bid evaluation matrix should be removed. If the vendor offers an unacceptable solution, then the vendor's proposal should be disqualified not down rated.

- Perhaps the procurement should be structured to ensure the vendor(s) has an adequate basis of experience. In particular, requiring prior delivery of a double crystal monochromator might prove to be a helpful filter for weeding out inexperienced vendors.

C.3 Recommendations

- It is recommended that the interface between the in-house design and a procurement contract be defined clearly prior to starting the advanced procurement.

- The specification and interface document should be revised prior to proceeding to advance procurement. In particular, the statement of work should be revised to limit the degrees of freedom that are the responsibility of the vendor consistent with the finding that many of the finer degrees of freedom can and should be included as part of the SLAC engineered crystal mounts.

RESPONSES TO THE REVIEW RECOMMENDATIONS

# XCS Diffractometer

# XCS : All the procurement documents regarding the XCS Diffractometer have been updated according to the findings and recommendations of the Review Committee of the Advanced Procurement Review. The new documents have been released and will be provided for the bidding process to the potential vendors.

# Each point requiring a precise response is also answered in the following.

A.2 Findings:

- The sample goniometer system is in a sufficiently developed state though there are some technical design issues noted below that should be addressed before procurement.

  - Specification of the performance of each stage in isolation does not properly account the errors introduced by coupling multiple stages. A more robust specification would tolerance the degrees of freedom at the sample interaction point.

XCS : The ESD was written with the intent that the stated motion specifications are applicable for a complete assembled diffractometer system, measured with respect to the sample (i.e. center of rotation). The vendor that quoted our system understood this point. The overall incremental errors of each motion and their associated performances, is in any case described as satisfying the “sphere of confusion”. Also the stability of each motion (translation and rotation) is specified in micron or degrees per hour.

  - By the same token, the specification should define the sample positioning needs and performance envelope not the engineering solution. The specification should allow the vendor as much freedom of engineering as feasible consistent with meeting these requirements. Thus, for example, the YY translation stage need not be specified as interleaved between the theta and chi/phi stages.

XCS : The documents have been updated in order to reflect that the vendor can feel free to propose an engineering solution that meet the specifications. Therefor the stack order of motion is just presented as a suggestion.

  - Cross coupling problem of the vertical stage YY to the tilting stages of χ and φ is anticipated. As noted above in the same context, stack-up position of these stages should be re-visited. The vendor may have better capability to build and select the stage stack-up position. The order of the stacking of the various motions should not be dictated, but left to the vendor who is designing to specifications. So, the ESD should specify that the vendor should propose their capability on the order of stack-up of the stages to achieve the motion stability requirements.

XCS : Same comment as A.2.9.b

  - The technical specifications for stage YY should address the motion straightness of trajectory.

XCS : The ESD now addresses the motion straightness of the YY stage.

- The XX_air motion is a good idea to move the detector. However, it would be advisable to make the table rest in fixed points on the floor. It would be advisable to think about installation cups recessed in the floor to provide repeatability (+/- 0.5 mm) of the diffractometer position after motion. In addition, the number of air-pads between the base and the floor will affect the overall system dynamics (high natural frequencies) and positioning stability. To meet the position drift and vibration requirements specified in the ESD, the interface design should be re-visited to ensure overall positioning stability and make the device more robust to the floor vibration.

XCS : This point will be discussed with the vendor during the design phase of the diffractometer,. That point is effectively crucial and therefore a comment regarding that point is now included in the ESD.

- The ESD does not include its minimum warranty period. Thus, for example, the ESD may include “the warranty will start the date of approval of its final acceptance test from SLAC for a period of one year“. The procurement document includes only “Proposed guarantee“.

XCS : The ESD is not the appropriate location to provide any details regarding any warranty details. However the  procurement specification document has been updated in order to include this comment.

- Synthetic granite material may be considered as an option due to its excellent damping quality.

XCS : The ESD has been modified and offer the possibility to provide any material that would meet satisfactorily the specifications.

- “True Stone Distributing” was a supplier of the granite in the APS nanoprobe project. The LUSI may consider this company as a subcontractor to the vendor.

XCS : This information will be provided during the design phase to the selected vendor.

- There is a typo in Section 7.8.2 of the ESD.

XCS : The ESD has been corrected.

- The vendor would benefit from as much hutch environmental information as available.

XCS : The ESD has been updated accordingly with a new section (c.f. Sec 7.8.10).

- General procurement plan is appropriately documented though the "not acceptable" classification in the bid evaluation matrix should be removed. If the vendor offers an unacceptable solution, then the vendor's proposal should be disqualified not down rated.

XCS : The documents have been updated accordingly.

A.3 Recommendations:

- The specification should allow the vendor as much freedom of engineering as feasible consistent with meeting the performance requirements.

XCS : This documents have been updated accordingly.

- The XCS diffractometer proceeds to advance procurement with minor changes to the specifications.

XCS : All the documents have been corrected/updated according to and following the recommendations and findings of the Review Committee. There fore XCS will proceed to advance procurement.

# B. XCS Large Angle Detector Mover

# XCS : All the procurement documents regarding the XCS Large Angle Detector Mover have been updated according to the findings and recommendations of the Review Committee of the Advanced Procurement Review. The new documents have been released and will be provided for the bidding process to the potential vendors.

# Each point requiring a precise response is also answered in the following.

B.2 Findings:

- It is difficult to imagine that this system warrants award to an offshore vendor. Surely there are US vendors for machine hardware that could meet the specification. The plan to separate the detector motion assembly and the flight tubes from the main mover is a good idea. This will help in expediting the main detector mover design and procurement, while the details of the flight path and the detector table designs are completed in house. The division of effort envisioned whereby SLAC staff takes responsibility for the flight tube, beam stop, etc is appropriate.

XCS : No comment required.

- The stored energy associated with the evacuated drift tube is substantial. The ES&H issues associated with this stored energy in the presence of a thin window should be addressed early in the design effort. The use of a large area, thin Beryllium window should be re-evaluated in this context. A thin foil of Kapton specified in the ESD is preferred.

XCS : The Physics Requirement Document has been corrected and no longer mention beryllium as a suggestion to be investigated for the exit window. The ESD already recommend the use of Kapton and therefore does not require to be modified.

- The large angle detector mover design is based on one currently in use at APS in sector 30. While there are many similarities, some of the differences are significant to affect the design for the LUSI instrument.

XCS : The specification of the Large Angle Detector Mover are in agreement with this comment. Whereas the concept is similar to the HERIX spectrometer, the various ranges, and levels of precision are different, as stated in all the documents.

- The load of the detector and the associated flight path is significantly small compared to the APS device.  This has the advantage of reduction in cost and will provide better stability for the mover.

XCS : No comment required.

- Renishaw’s linear encoder used for HERIX at Sector 30, APS, has a reliability problem so other options to measure the displacement should be considered. Laser interferometer is an option to be evaluated.

XCS : XCS was aware of this issue and agrees. This will be discussed in details with the selected vendor during the design phase.

- Hardware controlled speed limit is critical for the safe operation of a detector mover with large travel range.

XCS : XCS agrees. This will be discussed in details with the selected vendor during the design phase.

- X-ray sensors may need to be added to the beam stop to provide an interlock for the detector equipment safety.

XCS : XCS agrees and will investigate the realization of this “ smart beamstop”. A comment is added in the ESD in Sec 7.2.2.1.

- In the physics document (SP-391-001-33), the Large Angle Detector Stage (LADS) should be the Large Angle Detector Mover (LADM).

XCS : The Physics Requirement Document has been corrected accordingly.

- There is a typo in Section 4 in the ESD (SP-391-001-31). The reference document should be SP-391-001-33.

XCS : The ESD has been corrected.

- In Section 4.1 of the procurement specification (PS-391-001-75), “Diffractometer system” should be “Main Carrier Mover subsystem of the Larger Angle Detector Mover”. In Section 10-d) of the procurement specification, “SP-391-001-30” should be “SP-391-001-31”.

XCS : The Procurement Specification has been corrected.

- General procurement plan is appropriate though the "not acceptable" classification in the bid evaluation matrix should be removed. If the vendor offers an unacceptable solution, then the vendor's proposal should be disqualified not down rated.

XCS : The documents have been updated accordingly.

- The vendor would benefit from as much hutch environmental information as available.

XCS : The ESD contains now a section on hutch environmental requirements.

B.3 Recommendations

- It is recommended that the large angle detector mover carriage should proceed to advance procurement after defining the interface between the in-house design and the mover carriage.

XCS : All the documents have been corrected/updated according to and following the recommendations and findings of the Review Committee. There fore XCS will proceed to advance procurement.

- The large Beryllium window for the detection drift tube has serious consequences for the ES&H if it ruptures. So, a thin foil of Kapton is recommended for this application.

XCS : We agree. The recommendation is already answered in details in Findings.B.2.

# C. XCS Large Offset Monochromator

# XCS : All the procurement documents regarding the XCS Large Angle Detector Mover have will be updated according to the findings and recommendations of the Review Committee of the Advanced Procurement Review. The new documents will be released and will be provided for the bidding process to the potential vendors.

# Each point requiring a precise response is also answered in the following.

C.2 Findings

- The specification for the large offset monochromator for LUSI is based on APS design. While there are major similarities between the two designs, care has to be taken to revise the specifications for the LUSI instrument.

XCS : The specifications will be reviewed to assure the description is appropriate for the LUSI instrument.

- The plan to use a translation stage in the Z axis along the beam direction for both first and second crystals is an excellent match for the LCLS source.  The design as called for is a mirror image between the first and second crystal motions.  The design also has the advantage of providing a fixed reference point for a potential detector between the two crystals. This monochromator concept of moving both crystals symmetrically about a fixed point is quite advantageous given the absence of cooling required of the first crystal. This greatly eases the Z rail tolerance problem.

XCS : The concept of moving both crystals in order to meet the Z-rail stability requirements will be retained.

- The LUSI staff is planning to design the crystal mounts and will be trying out different crystals during the course of their commissioning. Based on the past experience at SSRL and at APS, it is advisable to define a clear interface point for the crystal mount.

XCS : The RFQ will be written to allow the vendor to quote the monochromator with or without the crystal mounts. The interface of the crystal mount to the coarse translation stages will be defined.

- A number of the fine motion degrees of freedom are rightfully the domain of the crystal mount engineering and should be extracted from the vendor statement of work. The vendor SOW would benefit from reduced scope to include only the two larger translations and large angle rotation for each crystal. The SOW should be clear that the vendor needs to provide the means to swap crystals. Since vendors may not be thinking outside the box, the SOW might include descriptions of several means to swap crystals (e.g., vertical translation, 180 deg rotation and horizontal translation, ...)

XCS : The SOW will be revised to separate the fine translations from the coarse ones. The fine translations will be part of the crystal mount. The crystal changing requirement with examples will be expanded upon.

- As noted above in a different context, the specification should allow the vendor as much freedom of engineering as feasible consistent with meeting the monochromator requirements.

- Entropy is real, so the monochromator stability is likely to be enhanced by including the minimum number of degrees of freedom necessary. Also the lighter the load carried by a degree of freedom, the higher the resonant frequency. Thus fine motions are best implemented in the crystal mount.

XCS : The statement of work will stipulate that the vendor is allowed to propose their own solution to the monochromator basic requirements.

- As noted above in the same context, the removal of some of the motions to the crystal mounts will help in expediting the main monochromator procurement. Every extra degree of freedom is a source of operational instability and will degrade the performance of the device. Based on experience it would be advisable to remove the H1, H2,  and  advisable motions from the main stage and make them part of the crystal mount.  These designs exist at SSRL and can be reused in minimum work.  The experience of the SSRL engineering staff should be valuable in the design of the crystal mount.  The interface between the crystal mount and the rest of the stages should be clearly defined in the ESD and its crystal mount design effort by SLAC engineer should be included in budget.

XCS : The SSRL crystal mount designs will be reviewed. The degrees of freedom that  the vendor is required to provide will be adjusted according to which degrees of freedom will be accommodated by the crystal mount itself.

- Some crystal motions, such as Y1, Y2, etc., could be covered by a SLAC-home-made crystal holder as well.

XCS :Modification of the SLAC existing design crystal mount in order to accommodate the Y-motion will be investigated and the specification will be revised accordingly.

- Care has to taken to clearly define the tolerances on all the stages and ask for a testing plan of the deliverables.

XCS : The statement of work will include the requirement of developing a testing plan for each degree of freedom at the time of the Final Design Review. Approval of the final design will not be granted without an appropriate testing plan.

- Radiation hardened optical encoders may be necessary for encoders inside the vacuum chamber.

XCS : The need for radiation hardened encoders inside the vacuum chamber will be expressed and evaluated during the preliminary design phase.

- Support structures of the monochromator should have overall alignment capability (wedge Jack and X and Y pusher etc.)

XCS : The alignment requirements of the monochromator were stated in the Engineering Specification Documents. This section will be reviewed to assure clarity.

- The vendor would benefit from as much hutch environmental information as available.

XCS : Environmental information will be provided to the vendor. The environment requirements given to facilities for the construction of building area housing the monochromator will be provided initially. Measured environmental information will be provided once the construction of this area will be completed.

- “Angle” should be removed from a Title of “XCS Large Angle offset Monochromator Statement of Work – Statement of Work” (SP-391-001-80 R0)

XCS : This will be corrected.

- In Section 3.2.2. of procurement document (PS-391-001-78), “SP-3914-000-94“ should be “SP-391-000-94”. In Section 3.2.3. “Main Carrier Mover subsystem” should be “Large Offset Monochromator”.

XCS : This will be corrected.

- A possible US vendor for the monochromator is as follows;

Physical Sciences Laboratory

University of Wisconsin-Madison

3725 Schneider Dr. Stoughton WI 53589-3098

Phone:	(608)877-2200, Fax: (608)877-2201

Email: ffeyzi@psl.wisc.edu

Contact: Farshid Feyzi

XCS : The above mentioned vendor will be added to the vendor list.

- The project should plan for a 24 month lead time following order placement. The vendor may deliver in 18 months and the crystal mounts should be engineered accordingly, but the definition of success should be 24 months.

XCS : The procurement specification will specify a preferred delivery date of 18months. The vendor will be encouraged to provide us with a realistic delivery schedule.

- General procurement plan is appropriate though the "not acceptable" classification in the bid evaluation matrix should be removed. If the vendor offers an unacceptable solution, then the vendor's proposal should be disqualified not down rated.

XCS : The “ not acceptable” classification will be revised in a way that it will be clear that this grade grounds for the disqualification of the offer.

- Perhaps the procurement should be structured to ensure the vendor(s) has an adequate basis of experience. In particular, requiring prior delivery of a double crystal monochromator might prove to be a helpful filter for weeding out inexperienced vendors.

XCS : The procurement specification will include “ prior delivery of a double crystal monochromator” as an element of the evaluation of vendors.

C.3 Recommendations

- It is recommended that the interface between the in-house design and a procurement contract be defined clearly prior to starting the advanced procurement.

XCS : The interface between the crystal mount scope of work and the main translation stage will be clearly defined prior to the release of the request for quotation.

- The specification and interface document should be revised prior to proceeding to advance procurement. In particular, the statement of work should be revised to limit the degrees of freedom that are the responsibility of the vendor consistent with the finding that many of the finer degrees of freedom can and should be included as part of the SLAC engineered crystal mounts.

XCS : The ESD and the Statement of work will be revised and an interface control document will be produced prior to proceeding with the procurement of the Large Offset Monochromator. The number of degrees of freedom that are the responsibility of the vendor designing the main translation stages in the monochromator will be clearly separated from the ones that are the responsibility of the SLAC crystal mount effort. In addition, an option for the vendor to offer to provide the crystal mount will be included in the RFQ package.

XCS : All the documents will be corrected/updated according to and following the recommendations and findings of the Review Committee. There fore XCS will proceed to advance procurement when all the corrected documents will be released.

XPP

# Detector Positioning System

Findings:

o  The robotic solution to detector positioning offers an extremely adaptable positioning system. This adaptability is essential to accommodate the demanding detector positioning envelope and avoid conflicts with sample positioning and pump laser systems as well as accommodate evolution in detector positioning requirements as the XPP user community expands and matures.

o  The demanding detector positioning envelope is appropriate given the diverse character of the XPP science and associated instrumentation requirements.

o  Separation of the detector positioning system from the sample manipulation system is warranted to minimize adverse influences of the detector on the sample positioning system.

o  The results of the robot testing SOW provide confidence that a commercial robot system can provide the required performance.

o  Two detector positioning system procurement strategies are plausible. (a) An integrated solution procurement strategy utilizes a single RFP for the entire system including robot, robot mount, control software, and protection systems. (b) A subsystem procurement strategy involves multiple RFP for the different subsystems of the detector positioning system.

o  The primary advantage of the integrated solution is a single vendor assumes responsibility for the entire system, so definition of performance requirements and assessment of vendor product performance is relatively straight forward. The primary concern is no one vendor may have the requisite mix of capabilities. The detector positioning engineering specification document provides a reasonably complete description of the detector positioning requirements subject to some qualification as noted below. As such, with some modification, it could serve as the requirements document for the integrated solution procurement.

o  The subsystem procurement strategy may allow a better match of vendor competency to subsystem requirements. A subsystem procurement strategy, however, demands clear definitions of subsystem performance requirements that avoid coupling of performance assessment between the various subsystems. Thus the documentation and specification burden on project staff is much more significant with this procurement strategy as compared to the integrated solution strategy. Additionally, more of the system integration effort falls to the project staff. As presently written the detector positioning engineering specification does not provide an adequate basis document for a subsystem procurement strategy.

The ESD is not intended to be a procurement document. It is intended to give overall complete system requirements. The numbers that have come from the robot measurements are consistent with an integrated system. IE: the measurements of the robot motion were good enough to adequately allow for additional tolerance stack-up from the support system, etc. We can, and will have to, address the precision / tolerance budget due to the various sources and explicitly allocate numbers to the various equipment / subsystems. This will be done when we have a more mature design and better measurements of the system / hutch. The additional issues of temperature and hutch dynamics and stability have to be dealt with one way or the other and cannot be addressed, independent of the procurement strategy, without additional measurements, which I am now planning on.

o  The software requirements, instrumentation and controls requirements, and machine protection and personnel protection system requirements sections of the engineering specification document sp-391-00-62 r0 are insufficiently developed. The review panel would have benefitted from inclusion of a software and controls expert.

o  Sole source procurement of a robot was not adequately justified. In particular, the performance differences between the Staubli RX160 robot and the ABB IRB 2400-16 robot are minor (ie., 10um in positioning reproducibility).

There are other issues with regard to specing the robot. Reach, acceptable mounting aspect (upright, upside down, 90 deg), relative lengths between joints, etc.

o  The robot mount SOW does not provide adequate information regarding hutch vibration spectra as well as longer term drift to assess the relative stability of various robot mount locations (eg., floor mount vs. roof mount).

This SOW is not intended to address anything with regard to motion precision, only robot work envelope and mounting location options. Stability, dynamics, precision, etc are separate issues.

o  Real time correction of detector position drift using optical tooling for detector position monitoring is difficult owing to the demanding detector positioning envelope and interrupted lines of sight associated with the XPP experimental equipment. Consequently such correction while conceptually plausible may prove practically infeasible. Certainly a very large contingency should be reserved for such a scheme if real time correction is a likely outcome of the procurement.

I, for one, do not foresee real time correction. We do foresee the ability to improve the nominal accuracy of the robot motion by improving the calibration of the robot. (the robot that was used for the motion measurements was not calibrated, it only had nominal, generic motion calibration - transfer functions). We can also use improved systems for robot / detector alignment and potentially recover some overall motion accuracy (not precision) if needed.

Recommendations:

o  Prior to issuing detector positioning system RFP(s), hutch environmental effects on detector stability need to be analyzed in particular as applies to a roof mounted solution. (agreed, I have this task in my “to do” list.)

o  If an integrated system procurement strategy is employed, then the engineering specification document should be embellished to define unambiguous software requirements, instrumentation and controls requirements, and machine and personnel protection requirements. Additionally, relevant environmental information such as hutch floor, wall, and roof vibration spectra and longer term stability as well as hutch temperature control should be provided.

Again, the ESD is not a procurement document. The “detector  motion ESD” is only intended to address the overall technical requirements, not the specifies of procurement or subsystem design. I think that most of these issues should be addressed in the subsequent procurement document(s)  or additional ESD’s for the robot subsystems.

o  If a subsystem procurement strategy is utilized, then the specification documentation requires substantial modification and re-review prior to release for procurement. Complete and self contained specifications must be developed for each subsystem of the detector positioning system prior to the procurement of any individual subsystem. The position tolerance budget established in detector positioning system engineering specification document sp-391-00-62 r0 must be apportioned amongst the subsystems. Performance requirements and delivery acceptance criteria for each subsystem must permit unambiguous assessment of subsystem performance despite potential performance coupling to other detector positioner subsystems. The revised documentation should address the software, instrumentation and control, machine and personnel protection issues noted above as well as the hutch environmental parameters. Re-review of the system specification should be accompanied by a proof of principle solution to the detector positioning system requirements which utilizes the results of the two robot SOW currently in process or complete.

Again the detector mover ESD is not intended to be the only document describing the overall system. “substantial modification” ? “Addition documentation is required”, which we are aware of and planning to complete as appropriate, is a more accurate way of saying it I believe.

I would resist changing the ESD to address topics that it is not intended for.  If one is to assume that the  robot, support structure and software are to be procured, all from outside vendors (worst case as far as required documentation is concerned) than I would, at a minimum, envision a procurement spec or SOW for each item, and possibly an ESD for the support structure (which would address the hutch dynamics and tolerance budget issues, as engineering design inputs for analysis) and possibly an ESD for the software (that would address safety system requirements, etc).

# Sample Goniometer System

Recommendations:

o  Specify the performance of each degree of freedom at sample interaction point for each of the two complete goniometer system configurations in the absence of the SLAC mounting base.

We will add a section / comment that specifically states that the performance criteria apply to completed goniometer assemblies (either configuration) and the measurement datum is at / from the sample goniometer center of rotation.

o  Add hutch environmental information to the engineering specification document. Add limit switch and hard stop requirements as applicable to the engineering specification document.

Limit switch and hard stop info will be added. As stated above, beyond specifying a nominal temp and tolerance, and maybe some other environmental conditions, FOR THE ACCEPTANCE TESTING, I do not think that adding environmental conditions of the hutch to this ESD is relevant or correct. I cannot think of ever seeing such info added as part of a requirement such as this. It would be appropriate as part of engineering / design inputs (like for the robot arm support structure).

o  Specify the SLAC mounting base requirements with appropriate consideration for the aggregate base plus goniometer system stability at the sample interaction point for each degree of freedom.

o  Perform acceptance testing of each degree of freedom at sample interaction point for each of the two complete goniometer system configurations in the absence of the SLAC mounting base.